Advanced Patient Monitoring &

Delirium Early Warning System

Detailed Implementation Plan

Introduction to IoT ¦ Team Members Teja Pudi (2025111021) ¦ Tushita Arora (2025111019) Shaurya Bansal (2025101118) ¦ Dhruva Anand (2025111009)

April 2025

# 1. Project Overview & Architecture

This document provides a step-by-step technical implementation plan for building the wrist-worn ICU delirium early warning system. The system uses two ESP32 microcontrollers, five physiological sensors, on-device signal processing, a multivariate coherence model, and a dual-display alert system.

## 1.1 System Architecture Summary

| Layer | Components / Description |
| --- | --- |
| Sensing Layer | MAX30102 (PPG/SpO2), AD8232 (ECG), MAX30205 (Temperature), ADXL335 (Accelerometer), GSR Electrodes |
| Processing Layer | ESP32 #1 (Sensor Hub): raw acquisition, HRV extraction, feature computation. ESP32 #2 (Analytics Hub): coherence model, alert engine, display controller |
| Communication Layer | SPI / I2C between sensors and ESP32 #1. Serial / ESP-NOW between the two ESP32s. Wi-Fi (ESP32 #2) to remote dashboard |
| Display Layer | Bedside OLED (SSD1306 128x64) for live vitals. Browser dashboard (HTML+JS) hosted on ESP32 #2 web server for 6-hour trends |
| Alert Layer | Three-level RGB LED + buzzer output. Green / Yellow / Red state driven by score trajectory, not threshold crossings |

## 1.2 Data Flow

Sensors → ESP32 #1 (raw ADC sampling at 100-250 Hz per signal) → feature vector (RMSSD, SDNN, QT interval, SpO2 micro-variance, temp delta, movement magnitude, GSR level) every 30 seconds → sent to ESP32 #2 → appended to 60-sample rolling window → coherence model outputs stability score (0-100) → score trajectory evaluated → alert state set → OLED + dashboard updated.

# 2. Hardware Components — Detailed Specifications & Wiring

## 2.1 Microcontrollers

| Role | Specification |
| --- | --- |
| ESP32 #1 — Sensor Hub | ESP32-WROOM-32 Dev Board. Handles all sensor I/O, ADC reads, PPG/ECG signal processing, feature extraction. Runs FreeRTOS with task-per-sensor approach. Communicates to ESP32 #2 via UART (115200 baud) or ESP-NOW. |
| ESP32 #2 — Analytics Hub | ESP32-WROOM-32 Dev Board. Receives feature vectors, maintains 30-min rolling window, runs coherence model, manages OLED, hosts Wi-Fi web server for dashboard, drives RGB LED + buzzer alert system. |

## 2.2 Sensor Details and Pin Mapping

### MAX30102 — PPG & SpO2 Sensor

| Pin (MAX30102) | Connect To | Purpose |
| --- | --- | --- |
| VIN | 3.3V (ESP32 #1) | Power supply |
| GND | GND | Ground |
| SDA | GPIO 21 (ESP32 #1) | I2C Data |
| SCL | GPIO 22 (ESP32 #1) | I2C Clock |
| INT | GPIO 34 (ESP32 #1) | Data-ready interrupt (optional) |

Library: sparkfun-MAX3010x or Adafruit MAX3010x. Default I2C address: 0x57. Configure in heart-rate mode for raw PPG waveform at 100 sps; enable SpO2 mode for parallel SpO2 sampling at 25 sps. The raw IR waveform is the primary data for HRV extraction.

### AD8232 — ECG Sensor

| Pin (AD8232) | Connect To | Purpose |
| --- | --- | --- |
| 3.3V | 3.3V (ESP32 #1) | Power supply |
| GND | GND | Ground |
| OUTPUT | GPIO 36 / ADC1_CH0 (ESP32 #1) | Analog ECG signal out |
| LO+ | GPIO 32 (ESP32 #1) | Lead-off detection + |
| LO- | GPIO 33 (ESP32 #1) | Lead-off detection - |
| SDN | GND (always on) | Shutdown pin — tie low to keep active |

Sampling: 250 Hz via ADC on ESP32 (use ADC1 channels only during Wi-Fi operation). Store in 512-sample circular buffer. Apply 0.5-40 Hz digital bandpass filter (FIR or biquad IIR) before R-peak detection. Use Pan-Tompkins algorithm or threshold-adaptive derivative method for R-peak detection. Compute RR intervals from successive R-peaks to extract RMSSD, SDNN, and QT interval.

### MAX30205 — Body Temperature Sensor

| Pin (MAX30205) | Connect To | Purpose |
| --- | --- | --- |
| VDD | 3.3V (ESP32 #1) | Power supply |
| GND | GND | Ground |
| SDA | GPIO 21 (shared I2C) | I2C Data — same bus as MAX30102 |
| SCL | GPIO 22 (shared I2C) | I2C Clock — same bus as MAX30102 |
| A0/A1/A2 | GND/GND/GND | Address bits — I2C address = 0x48 |

Accuracy: +/- 0.1 degrees Celsius. Sample every 60 seconds (no need for high frequency). Store a 24-sample (24-minute) rolling buffer to fit a circadian temperature curve. Expected healthy core temperature follows a cosine wave with ~0.5 degree amplitude across a 24-hour period; use a least-squares cosine fit (A*cos(2*pi*t/1440 + phi) + C) and report the residual as the circadian deviation score.

### ADXL335 — 3-Axis Analog Accelerometer

| Pin (ADXL335) | Connect To | Purpose |
| --- | --- | --- |
| VCC | 3.3V (ESP32 #1) | Power supply |
| GND | GND | Ground |
| XOUT | GPIO 35 / ADC1_CH7 | X-axis analog output |
| YOUT | GPIO 34 / ADC1_CH6 | Y-axis analog output (if INT not used here) |
| ZOUT | GPIO 39 / ADC1_CH3 | Z-axis analog output |

Sensitivity: 300 mV/g at 3.3V supply. Sample at 50 Hz. Compute resultant acceleration magnitude: mag = sqrt(ax^2 + ay^2 + az^2). Subtract gravity baseline (1g = ~300 mV). Compute RMS of magnitude over a 30-second window to get a restlessness score. High-pass filter at 0.5 Hz removes gravity component and isolates movement. Flag as agitated if 30-second RMS exceeds threshold (tune empirically, starting at ~0.15g).

### GSR (Galvanic Skin Response) Electrodes

| Connection | Connect To | Purpose |
| --- | --- | --- |
| Electrode 1 (index finger) | One leg of 10k voltage divider | Skin conductance sensing |
| Electrode 2 (middle finger) | GND | Return path |
| Divider midpoint | GPIO 32 / ADC1_CH4 (ESP32 #1) | ADC read of skin conductance voltage |
| Divider 3.3V leg | 3.3V via 10k resistor | Fixed reference voltage |

Build a simple voltage divider: 3.3V -- 10k resistor -- ADC pin -- skin -- GND. Skin resistance varies from ~10k Ohm (high arousal) to ~2M Ohm (baseline calm). Sample at 10 Hz. Smooth with a 5-sample moving average. Report both tonic (slow baseline) and phasic (rapid change) components. Phasic spikes of more than 0.2 V above baseline within a 5-second window indicate sympathetic activation.

## 2.3 Additional Hardware

| Component | Wiring & Notes |
| --- | --- |
| SSD1306 OLED 128x64 (I2C) | VCC=3.3V, GND=GND, SDA=GPIO 21 (ESP32 #2), SCL=GPIO 22 (ESP32 #2). Address 0x3C. Use Adafruit_SSD1306 library. Show: HR, SpO2, Temp, HRV, Stability Score, Alert level. |
| RGB LED (common cathode) | R=GPIO 25 (ESP32 #2) via 220 Ohm, G=GPIO 26 via 220 Ohm, B=GPIO 27 via 220 Ohm. PWM-driven. Green=normal, Yellow=caution, Red=alert. |
| Passive Buzzer | GPIO 14 (ESP32 #2) via 100 Ohm resistor. Use ledcWriteTone() for beep patterns: 1 short beep (yellow), 3 rapid beeps (red). |
| 3.7V LiPo Battery | Connect to JST connector on ESP32 dev board (if available) or to a TP4056 charge module. Target 2000 mAh for ~8 hours wearable runtime. Include power switch on LiPo positive rail. |
| ESP32 to ESP32 link | UART: TX1 (GPIO 17) of ESP32 #1 to RX2 (GPIO 16) of ESP32 #2. TX2 (GPIO 17) of ESP32 #2 to RX1 (GPIO 16) of ESP32 #1. Baud: 115200. Alternative: use ESP-NOW over Wi-Fi for wireless link. |

# 3. Firmware Implementation — ESP32 #1 (Sensor Hub)

## 3.1 Development Environment Setup

- Install VS Code + PlatformIO IDE extension.

- Create new PlatformIO project. Board: espressif32 / esp32dev. Framework: Arduino.

- Add libraries to platformio.ini:

platformio.ini dependencies - lib_deps = sparkfun/SparkFun MAX3010x Pulse and Proximity Sensor Library - adafruit/Adafruit MAX31865 -- NOT needed; use: sparkfun/SparkFun MAX30205 - electroniccats/mpu6050 -- NOT needed (using ADXL335 raw ADC) - knolleary/PubSubClient -- for MQTT if needed - bblanchon/ArduinoJson @ ^6.21

## 3.2 Task Architecture (FreeRTOS)

ESP32 #1 runs four FreeRTOS tasks pinned to Core 1 (leave Core 0 for Wi-Fi if needed):

| Task Name | Description & Timing |
| --- | --- |
| TaskPPG_ECG | Samples MAX30102 IR channel + AD8232 ADC simultaneously at 250 Hz. Uses a hardware timer ISR to trigger ADC reads and push to a ring buffer. Runs every 4 ms. |
| TaskHRV | Every 30 seconds, reads the PPG/ECG ring buffer. Detects R-peaks, computes inter-beat intervals, calculates RMSSD = sqrt(mean of squared successive differences), SDNN = std dev of NN intervals, QT interval approximation. |
| TaskSlowSensors | Reads MAX30205 (temperature), ADXL335 (3-axis), GSR every 10 seconds. Applies moving averages. Computes restlessness RMS and GSR phasic spike count. |
| TaskSend | Every 30 seconds, packages the latest feature vector as JSON and sends over UART to ESP32 #2. JSON format: {ts, hr, spo2_var, rmssd, sdnn, qt, temp, temp_delta, movement_rms, gsr_tonic, gsr_phasic_spikes} |

## 3.3 Signal Processing — Key Algorithms

### PPG Peak Detection & HRV Extraction

- Apply a 0.5-4 Hz bandpass filter to the raw IR signal (heart rate range).

- Detect peaks using an adaptive threshold: threshold = 0.6 * rolling max over last 5 seconds.

- Compute peak-to-peak intervals (RR intervals) in milliseconds.

- Remove artifacts: discard any RR interval less than 300 ms or greater than 2000 ms.

- RMSSD = sqrt( sum((RR[i+1] - RR[i])^2) / N ) over the 30-second window.

- SDNN = standard deviation of all valid RR intervals in the window.

### QT Interval Estimation

- Use the ECG signal from AD8232 for more precise R-peak detection (higher SNR than PPG).

- Detect T-wave end using the slope-threshold method: find the point where the post-R derivative returns to baseline.

- QTc (corrected QT) = QT / sqrt(RR in seconds) using Bazett's formula.

- Flag prolonged QTc if QTc greater than 450 ms in males or 470 ms in females (configure as parameter).

### SpO2 Micro-Variance Score

- Standard SpO2 = (AC_red / DC_red) / (AC_ir / DC_ir) * calibration constant.

- Micro-variance score: compute variance of successive SpO2 readings over 2-minute window.

- A healthy patient shows stable SpO2 (variance less than 0.5%). Rising variance (greater than 1.5%) indicates autonomic instability even before the mean drops below 95%.

- Report: mean SpO2, variance score (0-10 normalised), and a flag if variance is trending upward for 3 consecutive windows.

# 4. Firmware Implementation — ESP32 #2 (Analytics Hub)

## 4.1 Data Reception and Rolling Window

ESP32 #2 maintains a circular buffer of 60 feature vectors (30-second intervals = 30 minutes of history). On receiving a new JSON frame over UART:

- Parse JSON with ArduinoJson.

- Validate fields (reject if HR=0 or any field is NaN).

- Push to circular buffer, overwriting oldest entry when full.

- Trigger coherence model computation.

## 4.2 The Multivariate Coherence Model

The coherence model produces a single patient stability score from 0 to 100 (100 = fully stable, 0 = critical). It is NOT a threshold check; it is a trajectory and co-variation model.

### Step 1: Per-Signal Normalization

For each of the 7 features, maintain a 60-sample rolling mean and standard deviation. Normalize each value to a z-score: z = (x - mean) / std. Clip z-scores to the range [-3, +3].

### Step 2: Direction Scoring

| Feature | Delirium Direction (points toward instability) |
| --- | --- |
| RMSSD | Decreasing (reduced HRV = autonomic dysfunction) |
| SDNN | Decreasing (less variability = rigidity) |
| QTc interval | Increasing (prolongation = cardiac stress) |
| SpO2 variance | Increasing (instability in oxygenation) |
| Temp deviation from circadian fit | Increasing (circadian rhythm flattening) |
| Movement RMS | Increasing (restlessness, agitation) |
| GSR phasic spikes per window | Increasing (sympathetic activation) |

### Step 3: Weighted Coherence Score

Assign weights based on clinical significance: RMSSD (0.25), SDNN (0.15), QTc (0.20), SpO2 variance (0.15), Temp delta (0.10), Movement (0.10), GSR (0.05). Total = 1.00.

Instability score = sum(weight_i * normalized_direction_i) for all features moving in the delirium direction.

Stability score = 100 * (1 - instability_score / max_possible_instability).

### Step 4: Trajectory Evaluation (Alert Logic)

- GREEN (score >= 70 and stable or improving): normal monitoring, 30-second update rate.

- YELLOW (score 50-69 OR score fell more than 10 points in 10 minutes): early warning. Increase OLED refresh to 10 seconds. Send push notification to dashboard.

- RED (score < 50 OR score fell more than 20 points in 15 minutes): delirium alert. Activate buzzer. Flash red LED. Log timestamped alert event to SPIFFS.

The key innovation: a score of 65 that was 75 ten minutes ago triggers YELLOW even if 65 alone would be green. Trajectory beats instantaneous value.

## 4.3 OLED Display Layout (128x64)

OLED Screen Layout (SSD1306 128x64) - Row 1 (y=0): HR: 72 bpm SpO2: 97.2% - Row 2 (y=16): Temp: 37.1C HRV: 42ms - Row 3 (y=32): Stability: [===------] 58 - Row 4 (y=48): STATUS: YELLOW 12:34:07

Use Adafruit_SSD1306 + Adafruit_GFX. Refresh display every 10 seconds in yellow/red states, every 30 seconds in green. Draw a horizontal bar for the stability score using drawRect() and fillRect(). Invert display background in RED state for visual urgency.

## 4.4 Web Dashboard (ESP32 #2 Web Server)

- ESP32 #2 connects to hospital Wi-Fi (or creates its own AP if no network available).

- Runs an AsyncWebServer on port 80. Serves a single-page HTML dashboard stored in SPIFFS.

- Exposes a /data endpoint returning the last 720 data points (6 hours) as JSON.

- Dashboard HTML uses Chart.js (CDN) to render 4 time-series charts: stability score, HR+HRV, SpO2 variance, Temp deviation.

- Auto-refreshes every 30 seconds via JavaScript fetch(). Color-codes chart background red/yellow based on current alert state.

- Includes a scrollable event log table showing all state transitions with timestamps.

# 5. Software Implementation — Dashboard & Storage

## 5.1 SPIFFS File System (ESP32 #2)

- Format SPIFFS partition in PlatformIO: set board_build.partitions = min_spiffs.csv in platformio.ini.

- Store dashboard HTML/JS/CSS as static files in the /data folder (PlatformIO uploads with pio run --target uploadfs).

- Log alert events to /events.log in CSV format: timestamp, score, state, trigger_reason.

- Rotate log file when it exceeds 50 KB (overwrite oldest 20% of entries).

## 5.2 Dashboard HTML/JS Structure

| Component | Implementation |
| --- | --- |
| Score Trend Chart | Chart.js line chart. X-axis: time (last 6 hours). Y-axis: 0-100. Three colored zones: green fill (70-100), yellow fill (50-70), red fill (0-50). Current score shown as large number. |
| HR & HRV Chart | Dual-axis Chart.js line chart. HR (bpm) on left axis. RMSSD (ms) on right axis. Overlay shows R-peak artifacts flagged by the firmware. |
| SpO2 Variance Chart | Line chart. Shows mean SpO2 and variance band. Highlights windows where variance exceeded 1.5%. |
| Temperature Chart | Shows raw temperature + fitted circadian cosine curve. Deviation from expected curve plotted in red. |
| Alert Events Table | Scrollable table. Columns: Time, Event, Score Before, Score After, Duration. Color-coded rows by severity. |
| Patient Header | Static div at top: Patient ID, Room, Sensor battery %, last update time, current alert badge. |

## 5.3 Inter-ESP32 Communication Protocol

JSON Feature Vector Format (ESP32 #1 -> ESP32 #2, every 30 sec) - {"ts":1712345678, "hr":72, "rmssd":38.4, "sdnn":52.1, "qtc":435, - "spo2_mean":97.2, "spo2_var":0.8, "temp":37.12, "temp_delta":0.15, - "movement_rms":0.08, "gsr_tonic":3.2, "gsr_spikes":2, "lead_off":false}

Baud rate: 115200. Framing: newline-delimited JSON (one JSON object per line). ESP32 #2 reads lines via Serial2.readStringUntil('\n'). Watchdog: if no frame received for 90 seconds, trigger a YELLOW alert and display 'SENSOR TIMEOUT' on OLED.

# 6. Phased Build & Testing Plan

Follow these phases sequentially. Do not advance until all tests in the current phase pass.

PHASE 1 Hardware Bringup & Individual Sensor Validation

- Power each sensor individually from ESP32 3.3V rail. Measure current draw with multimeter.

- MAX30102: Run SparkFun library example. Verify red and IR values change when finger is placed. Expected raw IR: 50,000-100,000 counts.

- AD8232: Connect 3-lead electrodes. Run serial plotter. Verify visible QRS complexes at rest. P and T waves optional at this stage.

- MAX30205: Run I2C scanner sketch. Verify address 0x48 is detected. Read temperature. Should read body temperature when held.

- ADXL335: Read all 3 ADC channels. At rest flat, Z should read ~512 (1g = 3.3V/2 + 300mV offset). Rotate sensor and verify axis values shift.

- GSR: Build voltage divider. Wet fingers = lower resistance = higher ADC reading. Dry fingers = lower ADC reading. Verify this relationship.

- PASS CRITERIA: All 5 sensors return plausible data in isolation on the benchtop.

PHASE 2 ESP32 #1 Firmware — Sensor Fusion & Feature Extraction

- Implement TaskPPG_ECG with 250 Hz timer ISR. Use oscilloscope or logic analyser to verify ADC sampling rate.

- Implement R-peak detection. Plot detected peaks against raw signal in Serial Plotter. Validate peak timing is correct.

- Compute RMSSD and SDNN on a known simulated dataset (pre-store a 30-second RR interval array and verify against manual calculation).

- Implement TaskSlowSensors. Run for 10 minutes. Log all features to Serial. Check for NaN or obviously wrong values.

- Implement TaskSend. Parse the JSON on a PC using Python (json.loads) to confirm valid format.

- PASS CRITERIA: Feature vector sent every 30 seconds. RMSSD within +/- 2 ms of manually computed value on test dataset.

PHASE 3 ESP32 #2 Firmware — Coherence Model & Alerts

- Stub out the UART receiver with hardcoded test feature vectors (simulate 60 frames: 30 normal, then 30 progressively degrading).

- Implement rolling window and normalization. Verify z-scores remain bounded after warm-up period (first 10 frames).

- Implement coherence model. With normal data, score should be >= 75. With all signals trending toward delirium, score should fall below 50 within 10 minutes.

- Connect real ESP32 #1 and verify score responds to deliberate sensor perturbation (briefly covering MAX30102 with fingers to distort SpO2, gentle tapping to raise movement score).

- Test alert transitions: green -> yellow -> red. Verify LED color changes and buzzer sounds correct pattern.

- PASS CRITERIA: Score responds to sensor input within two 30-second windows. All three alert levels trigger correctly with correct LED and buzzer output.

PHASE 4 Display & Dashboard Integration

- Implement OLED display. Verify all four rows render correctly with live data.

- Upload dashboard HTML to SPIFFS. Connect to ESP32 #2 Wi-Fi AP. Verify page loads on a phone browser.

- Run system for 30 minutes. Verify 6-hour trend charts populate correctly (scale x-axis to available data when less than 6 hours of data exists).

- Test auto-refresh: data on dashboard should update every 30 seconds without page reload.

- PASS CRITERIA: Dashboard loads in less than 3 seconds. Charts show correct time-aligned data. Alert events appear in the event log.

PHASE 5 Wearable Integration & Extended Testing

- Assemble on wrist mount. Verify sensors maintain contact during mild movement (typing, arm rotation).

- Run continuous monitoring for 2 hours on a healthy subject. Log all data. Inspect for signal dropouts and spurious alerts.

- Induce a simulated alert: exercise briefly to increase movement + HR, then immediately rest (simulates agitation followed by sudden stillness). Verify yellow alert triggers.

- Battery life test: run continuously from full charge. Record time to 20% battery (target: > 7 hours).

- PASS CRITERIA: Less than 5% packet loss over 2 hours. No spurious RED alerts in 2-hour healthy subject test. Battery > 7 hours.

# 7. Calibration & Tuning Guide

## 7.1 Per-Patient Baseline Calibration

On first application (or after any sensor re-attachment), run a 10-minute baseline period. During baseline:

- All features are collected but no alerts are triggered (grace period).

- The rolling window warm-up initializes the normalization statistics.

- A 'personal baseline' snapshot is saved to SPIFFS: baseline_HR, baseline_RMSSD, baseline_temp, baseline_GSR.

- All subsequent z-scores are computed relative to this personal baseline rather than a population average.

This eliminates false positives due to inter-patient variation (e.g., patients with chronic low HRV or naturally elevated temperature).

## 7.2 Tunable Parameters

| Parameter | Default Value | Tuning Guidance |
| --- | --- | --- |
| Alert YELLOW threshold | Score < 65 OR drop > 10 in 10 min | Lower threshold in high-risk patients; raise for less sensitive wards. |
| Alert RED threshold | Score < 50 OR drop > 20 in 15 min | Tune based on false positive rate in first week of deployment. |
| RMSSD weight | 0.25 | Increase to 0.30 if ECG lead quality is consistently high. |
| QTc prolongation flag threshold | 450 ms (male), 470 ms (female) | Set per patient based on their baseline QTc from admission 12-lead ECG. |
| SpO2 variance alarm threshold | Variance > 1.5% | Raise to 2.0% in patients with known sleep apnea. |
| Movement agitation threshold | RMS > 0.15 g | Lower to 0.10 g for sedated ICU patients with expected low movement. |
| Circadian model window | 24 hours | Not tunable at this stage; future enhancement for shifted circadian rhythms. |

# 8. Bill of Materials & Estimated Cost

| Component | Quantity | Estimated Cost (USD) |
| --- | --- | --- |
| ESP32-WROOM-32 Dev Board | 2 | $8 - $12 each |
| MAX30102 Pulse Oximeter Module | 1 | $4 - $7 |
| AD8232 ECG Module (with electrodes) | 1 | $8 - $14 |
| MAX30205 Temperature Sensor Breakout | 1 | $3 - $6 |
| ADXL335 Accelerometer Module | 1 | $3 - $6 |
| GSR Electrodes (metal snap or gel) | 2 | $1 - $3 |
| SSD1306 128x64 OLED Module (I2C) | 1 | $3 - $6 |
| RGB LED (common cathode) | 1 | $0.10 |
| Passive Buzzer (5V compatible) | 1 | $0.50 |
| 3.7V LiPo Battery (2000 mAh) | 1 | $6 - $10 |
| TP4056 LiPo Charge Module | 1 | $0.50 |
| Resistors (220 Ohm, 10k, 100 Ohm) | Assorted | $1 |
| Wristband / Flexible PCB mount | 1 | $2 - $5 |
| Breadboard + jumper wires | 1 set | $3 - $5 |
| TOTAL ESTIMATED |  | $43 - $77 |

# 9. Risks, Mitigations & Known Challenges

| Risk / Challenge | Mitigation Strategy |
| --- | --- |
| ECG lead-off on wrist-worn device (wrist ECG has poor signal quality compared to chest leads) | Use AD8232 in 2-lead configuration across forearm. Apply conductive gel. If SNR is too low, fall back to PPG-only HRV and reduce QT interval weight in the coherence model. |
| Motion artifact corrupting PPG and ECG simultaneously | Cross-check: if both PPG and ADXL335 show high-movement simultaneously, flag the window as artifactual and exclude it from the coherence model. Do not decrease score on artifact windows. |
| ESP32 ADC non-linearity (ESP32 ADC is known for non-linearity near 0V and 3.3V) | Keep all sensor outputs in the 0.3V - 2.5V range. Apply the ADC linearization lookup table published by Espressif (esp_adc_cal library). |
| I2C bus conflicts between MAX30102 and MAX30205 | Both are on the same I2C bus. Use Wire.setClock(100000) for reliable operation. Test with I2C scanner to confirm both addresses (0x57 and 0x48) resolve independently. |
| Rolling window warm-up period producing unreliable early scores | Do not display or alert on scores until at least 10 feature vectors (5 minutes) have been collected. Show 'Initializing...' on OLED during warm-up. |
| Wi-Fi interfering with ESP32 ADC (ADC2 channels are unavailable during Wi-Fi) | Use ONLY ADC1 channels on both ESP32s. All sensor pins must be on GPIO 32-39 for ADC functionality when Wi-Fi is enabled on ESP32 #2. |
| Battery life shorter than expected due to continuous Wi-Fi | Put ESP32 #2 Wi-Fi in light-sleep mode between dashboard refresh windows. Transmit only on client request. Estimate Wi-Fi active duty cycle at ~5%. |

# 10. Suggested Team Assignments

| Team Member | Primary Responsibilities |
| --- | --- |
| Teja Pudi (2025111021) | ESP32 #1 firmware: PPG/ECG signal acquisition, R-peak detection, RMSSD/SDNN/QTc computation. Phase 1 and Phase 2 lead. |
| Tushita Arora (2025111019) | ESP32 #1 firmware: Temperature circadian model, ADXL335 movement scoring, GSR feature extraction. Phase 2 support and calibration. |
| Shaurya Bansal (2025101118) | ESP32 #2 firmware: Coherence model, rolling window, alert logic. Phase 3 lead. Tuning and threshold optimization. |
| Dhruva Anand (2025111009) | ESP32 #2 firmware: OLED display, web dashboard (HTML/JS/Chart.js), SPIFFS storage, UART protocol. Phase 4 lead. Integration testing. |

All members share responsibility for Phase 5 (wearable integration and extended testing). Weekly syncs recommended to align on JSON schema and alert threshold decisions.